C64 70 MHz SAW Filter Bank

Test Procedure

The unit after assembly is to be tested for functionality using the polyscope.

This will require following apparatus:

- Polyscope

- RG 223 RF cable Type-N female to TNC female

- RG 223 RF cable Type-N male to TNC male

- +19 V and +/- 12 V power supply

The experimental setup should be as shown in figure 1. Set the Ployscope to span 32 MHz around 70 MHz centre frequency. With RF power as -20 dBm, and vertical scale as 10 dBm/division , zero the polyscope. Now select the 6MHz bandwidth and ensure that the insertion gain at 70MHz is approximately 10 dB. Rejection at 55MHz and 85MHz should be at least -40dBc. Now select the 16MHz bandwidth. The insertion gain at 70MHz should be approximately 5dB. Rejection at 55MHz and 85MHz should be at least -40dBc. Now select the through path. Insertion gain should be approximately 2dB throughout the band. Note down the above observations in the appropriate columns provided in table 1.

After this test, the unit has to be kept powered on in an air-conditioned test laboratory for ten days from a +19V power supply(burn in qualifier test).

After the burn test is complete, the functionality test that was performed before the burn in test, has to be repeated, and the results noted down in the appropriate columns in the table. The results again lie in the specified limits.

The unit is now ready for final Qualification and Characterization. This will now require the following apparatus:

- HP 8753c network analyser, along with the HP 85046A S-parameter test and HP 8503B calibration test

- HP colorpro plotter at a proper HPIB address (set as per HP 8753C)

- RG 223 RF cable Type-N male to TNC female

- RG 223 RF cable Type-N male to TNC male

- +19 V and +/-12 V power supply

Calibrate the network analyser as explained in HP 8753C network analysers' manual. The procedure and the setup required to measure the insertion loss is explained at length in the same manual and will not be explained here. (A copy of the relevant pages of the manual is attached at the end of the write up for C63 for users' convenience.)While calibrating, set the RF stimulus to -20dBm. Follow the necessary steps to get a display resembling plot 1. The insertion loss and rejection should be within the criterion already specified in table 1. Take a hard copy to resemble the one shown in plot 1 by superimposing the three plots corresponding to the three selections on the same paper loaded onto the plotter (It is essential not to remove the paper unless all the three plots have been superimposed.) The grid, markers and text should be enabled only for one of the three plots. While superposing the other plots, only trace should be plotted and the other selections should be turned off by pressing necessary key.